영	문	규	격	서

COMMODITY DESCRIPTION

| 관세분류번호 | 품목번호 | 품 목 및 규 격 | 단 위 | 수 량 |
| --- | --- | --- | --- | --- |
| HS No. | Item No. | Description | Unit | Quantity |
| 9027.30.2000 |  | UV/VIS Spectrophotometer | set | 1 |

A. Features :

- A high-performance Seya-Namioka type monochromatic light device using high- resolution aberration correction Mechanical Ruled Concave Diffraction Grating. It has excellent linearity when preparing a calibration curve for high-concentration samples with low stray light, it is possible to quantitative analysis .

- Because the sample chamber is designed with a completely symmetrical optical system, the beam size is always constant regardless of the sample position and slit width change, so high-performance measurement is possible during sample analysis

- Baseline can be memorized in 3 types, so baseline instability that occurs when measuring different types of samples can be eliminated, so the reproducibility of output data is improved.

- Even for highly reactive samples, it is easy to measure because it can move quickly up to 2,400 nm/min at a low speed, and the wavelength can move up to 4,500 nm/mi when using the Go to λ function.

- Automatic Wavelength Calibration function and Self-Diagnosis Function/Auto start function are built-in, so it is possible to reduce the cost and time required for system inspection during sample analysis and to use the system very conveniently.

| B. | Specifications : |  |
| --- | --- | --- |
|  | 1. | Monochromator : |
|  |  | Grating-grating Double monochromator Type and High performance |
|  |  | Seya-Namioka mount |
|  | 2. | Wavelength range : 190 to 900nm |
|  | 3. | Resolution : 0.1nm |
|  | 4. | Bandpass : 0.1, 0.5, 1, 2, 4, 5nm (settable in 6 steps) |
|  | 5. | Stray light : 0.015%T (Nal 220nm, NaNO2 340nm) |
|  | 6. | Wavelength accuracy : |
|  |  | ±0.1nm (automatic wavelength calibrating function built in) |
|  | 7. | Wavelength setting reproducibility : |
|  |  | ±0.05nm |
|  | 8. | Photometric mode : Abs, %T, %R, E(S), E(R) |
|  | 9. | Photometric range : |
|  |  | Abs : -3.8 to 3.8 (effective range) |
|  |  | %T : 0 to 300 %T |
|  | 10. | Photometric accuracy (checked with NISTSRM 930 filter) |
|  |  | ±0.002 Abs (0 to 0.5 Abs) |
|  |  | ±0.003 Abs (0.5 to 1.0 Abs) |
|  |  | ±0.006 Abs (1.0 to 2.0 Abs) |
|  |  | ±0.3 %T |
|  | 11. | Photometric repeatability (checked with NISTSRM 930 filter) |
|  |  | ±0.001 Abs (0 to 0.5 Abs) |
|  |  | ±0.0015 Abs (0.5 to 1.0 Abs) |
|  |  | ±0.003 Abs (1.0 to 2.0 Abs) |
|  |  | ±0.1 %T |

| 12. | Baseline flatness : Within ±0.0003 Abs (190~850nm) |  |  |
| --- | --- | --- | --- |
| 13. | Baseline stability : Within 0.0002 Abs/h (at 500nm) |  |  |
| 14. | Baseline memory : 3 channels (System : 1 ch, User : 2 ch) |  |  |
| 15. | Wavelength scan speed : |  |  |
|  | 1.5, 3, 15, 30, 60, 120, 300, 600, 1,200, 1,800, 2,400 nm/min, |  |  |
|  | automatically controlled ; |  |  |
|  | Go to λ : 4,500 nm/min |  |  |
| 16. | Light source : |  |  |
|  | Adjustment-free deuterium lamp, (UV range) |  |  |
|  | Adjustment-free iodine tungsten lamp (50W, Visible range) |  |  |
| 17. | Light source switching : |  |  |
|  | Automatic switching interlocked with wavelength scan |  |  |
| Switching wavelength optionally selectable within a range of 325 to 370 nm |  |  |  |
| 18. | Sample compartment : |  |  |
|  | Beam interval : 100mm |  |  |
|  | 120(W) x 300(D) x 140(H) mm |  |  |
| 19. | Detector : Photomultiplier |  |  |
| 20. | Power requirement : AC 220 Volt, 60Hz |  |  |
| 21. | Data processor |  |  |
|  | 1) Data system | 1 | system |
| Standard Equipment : |  |  |  |
| o | U-3900 Main Spectrophotometer with Two cell holder | 1 | set |
| o | Tool kit (screwdrivers, etc.) | 1 | set |
| o | Instruction manual | 1 | copy |
| o | Standard operation software for U-3900 | 1 | ea |

C. Accessories :

- 60 INTEGRATING SPHERE ACCY	1 ea

- Glass Filter Holder ACCY	1 ea

- 5 DEGREE SPECULAR REFRECTANCE ACCY.	1 ea

D. Remarks :

- Warranty : 1 year after Installation.

- Installation & Operation :

Will be executed by factory trained engineer free of charge.